Method Validation Report No: MVL-524

| Active Ingredient | Avanafil | Test | Assay of Avanafil |
| --- | --- | --- | --- |
| Testing Procedure | RD-AVN-SM | SOP Followed | 3/QC/GEN/012 |
| Department | Research & Development | Revision No | 00 |

| Personnel’s Involved in Validation Study | Signature |
| --- | --- |
| Prepared By: R&D Analyst |  |
| Reviewed By: SAS Manager |  |
| Approved By: Sr.GM Quality Operations |  |
| Date of Approval: |  |

(Not Valid without Signature)

Pre-VERIFICATION Requirements - Installation Qualification (IQ) Documented Validation that all key aspects of an installed High performance Liquid chromatography system (Shimadzu) adhere to the approved design specification and that the recommendations of the Shimadzu have been suitably considered. - Operational Qualification (OQ) Operational Validation carried out after installation that shows High performance Liquid chromatography system Shimadzu Prominence-I LC-2030 Plus performs in accordance with Shimadzu specifications and process requirements and that the appropriate GMP systems (e.g. training, calibration, and maintenance, etc.) are in place. - Calibration status of equipment Equipment was calibrated and bears calibration sticker of the external calibrator. - Facilities: The Validation of the method of Determination was carried out in the Analytical Lab. of Shaigan Pharmaceuticals (PVT) LTD. in Rawalpindi, PAKISTAN. - Identification of Machine/ Equipment used: HPLC QC-196 - Traceability of Material or Product used for study: - Avanafil Working Standard # CPL/WS/25/AVN/038 - Avanafil Starting Material B # 15048002-AVN - Apparatus/Glassware used: - 50ml volumetric flask type “A” - 25ml volumetric flask type “A” - 1ml pipette type “A” - 1000 ml cylinder type “A” - 1000ml beaker type “A” Precautions: - Use Pyrex type “A” glassware and broken glassware should not be used. - During analysis temperature should be in between 20-25°C. - Use dried glassware. - Use HPLC grade solvents only and must be filtered with 0.45µ Filter - Always degas the mobile phase, as air bubbles may tend to form during solvent mixing or during temperature or pressure changes. Air bubbles may cause pump malfunctions and detector signal noise. - Never run the column dry. Make sure there is enough solvent in the reservoir. - Use a 0.2µ sized syringe filter for sample. Always filter your sample before injection.

Method of Analysis to be Validate. Column: Elution: Flow Rate: Detection: Injection Volume: Column Temp: Solution A: Solution B: Mobile Phase: Table 1 Time (Minutes); Solution A (%); Solution B (%). 0. 2. 22. 27. 30. 35. Standard solution: Sample Preparation: Procedure for Injection: Calculation: Result = AT/AS x Cs / CT x P Where, AT: Area of sample solution AS: Area of standard solution CS: Concentration of standard solution (mg/ml) CT: Concentration of sample solution (mg/ml) P: Purity of working standard Method Specification: Manufacturer’s Specification

Parameters Under Study

1. Linearity and Range 2. Accuracy and recovery 3. Precision 3.1. Repeatability 3.2. Intermediate Precision (Ruggedness) 3.2.1 Within Days Variation 3.2.2 By different analyst 4. Robustness 4.1 Changing Mobile Phase Composition 5. Force Degradation 6.Detection Limit 7.Quantitation Limit 8. Specificity

| 1. Linearity and Range | For the establishment of linearity concentration ranges of a test substance (from 80% to 120 %) were prepared to determine assay of drug substance. The range of this analytical method lies between the 80% to 120% of the test concentration |
| --- | --- |

| Concentration % | Conc. mg/ml | Peak area; where n=2 |
| --- | --- | --- |
| 80 | 0.122 | 8032483 |
| 90 | 0.135 | 9103237 |
| 100 | 0.150 | 10122453 |
| 110 | 0.165 | 11136604 |
| 120 | 0.182 | 12235850 |

Graphical Presentation

Correlation Coefficient

| Acceptance Criteria | Result |
| --- | --- |
| Correlation Coefficient≥ 0.997 | 0.999 |

| 2.Accuracy and Recovery | In the study of accuracy / recovery, Known amount of sample corresponding to three concentration levels i.e. 80%, 100 % and 120 % were taken. Results were obtained using three concentration levels (80%, 100%, and 120 %) and three replicate of each concentration, 80%is (lowest concentration) and 120% (highest concentration) of the expected working range. Assessment of accuracy was established by evaluating the recovery of the analyte (Percentage recovery) across the range of assay. Acceptance criteria: (theoretical amount ± 2%) |
| --- | --- |

# Accuracy and Recovery

Sample I:

| Level % | Conc.mg/ml | Peak Area | %RSD | %age Recovery |
| --- | --- | --- | --- | --- |
| 80 | 0.122 | 8181863 | 0.26 | 80.50 |
|  |  | 8223041 |  | 80.90 |
|  |  | 8194699 |  | 80.62 |

Sample II:

| Level % | Conc.mg/ml | Peak Area | %RSD | %age Recovery |
| --- | --- | --- | --- | --- |
| 100 | 0.150 | 10202900 | 0.78 | 100.38 |
|  |  | 10169861 |  | 100.06 |
|  |  | 10052852 |  | 98.90 |

Sample III:

| Level % | Conc.mg/ml | Peak Area | %RSD | %age Recovery |
| --- | --- | --- | --- | --- |
| 120 | 0.182 | 12098125 | 0.11 | 119.03 |
|  |  | 12106604 |  | 119.11 |
|  |  | 12081584 |  | 118.86 |

Grand average %age recovery of three Samples:

| Acceptance Criteria ± 2% | Average % Recovery |
| --- | --- |
| For 80 % Level 78 – 82 % | 80.67% |
| For 100 % Level 98 – 102 % | 99.78% |
| For 120 % Level 118 – 122 % | 119.00% |

| 3. Precision (Repeatability) | Peak areas of Six replicates of homogenous sample were determined. Mean value, Standard deviation and relative standard were calculated. The relative standard deviation of six determinations at 100% of the test concentration should not be greater than 2% for drug product. |
| --- | --- |

# 3.1 Precision (Repeatability)

| # Replicate | # Peak Area |
| --- | --- |
| Replicate 1 | 10120780 |
| Replicate 2 | 10042552 |
| Replicate 3 | 10034339 |
| Replicate 4 | 10070220 |
| Replicate 5 | 10071068 |
| Replicate 6 | 10094898 |
| # Mean | 10072310 |
| # Standard Deviation | 32226.71 |
| # Relative Standard deviation (RSD%) | 0.32 |

| Acceptance Criteria | Result |
| --- | --- |
| RSD < 2.0 % | 0.32 |

| # 3.2. # Intermediate Precision (Ruggedness) | Intermediate precision (Ruggedness) expresses within laboratory variation as on different days and with different analyst. |
| --- | --- |

# 3.2.1 Within Days Variation

| Day | Sample Area | Assay% | Average Assay % | %RSD |
| --- | --- | --- | --- | --- |
| Day 1 | 10054095 | 99.48 | 100.53 | 0.99 |
|  | 10171758 | 100.64 |  |  |
|  | 10254155 | 101.46 |  |  |
| Day 2 | 10238001 | 100.73 | 101.29 | 0.49 |
|  | 10330471 | 101.64 |  |  |
|  | 10317825 | 101.51 |  |  |

| Acceptance Criteria | Results% |
| --- | --- |
| RSD < 2.0 % | 0.99 |
|  | 0.49 |

# 3.2.2. By Different Analyst

| Analyst | Peak Area | Assay % | Average Assay % | %RSD |
| --- | --- | --- | --- | --- |
| 1 | 10054095 | 99.48 | 100.53 | 0.99 |
|  | 10171758 | 100.64 |  |  |
|  | 10254155 | 101.46 |  |  |
| 2 | 10273395 | 101.65 | 101.59 | 0.07 |
|  | 10259195 | 101.51 |  |  |
|  | 10268662 | 101.60 |  |  |

| Acceptance Criteria | Result % |
| --- | --- |
| RSD < 2.0 % | 0.99 |
|  | 0.07 |

| # 4. ROBUSTNESS | Robustness of an analytical procedure is a measure of its capacity to remain unaffected by small but deliberate variations in procedural parameters. (Acceptance criteria: Method should be robust and it should not show a change in assay from the actual method parameters by more than 2%) |
| --- | --- |

| Mobile Phase Flow Rate | Peak Area |  |  | Assay % |
| --- | --- | --- | --- | --- |
| 0.4ml/min | 10250991 |  |  | 101.35 |
|  | 10424630 |  |  |  |
| 0.6ml/min | 10265336 |  |  | 101.48 |
|  | 10452302 |  |  |  |

| # 5. Forced Degradation | Forced degradation is a measure to analyze the effect of different solvents on the active pharmaceutical ingredient. Sample Solutions were subjected to stress conditions in 0.1N HCl, 0.1N NaOH and Hydrogen Peroxide Solution 3%. The chromatographic analysis of the sample solutions were carried out under same conditions. Basic Degradation: Add 5ml of 0.1N NaOH solution to the standard solution and make up volume with diluent. Acid Degradation: Add 5ml of 0.1N HCl solution to the standard solution and make up volume with diluent. H2O2 Degradation: Add 5ml of 3% H2O2 solution to the standard solution and make up volume with diluent. |
| --- | --- |

| Degradation Type | Peak Area | Assay % | Average Assay % |
| --- | --- | --- | --- |
| Acid Degradation (0.1N HCl) | 10193883 | 100.29 | 100.27 |
|  | 10189922 | 100.25 |  |
| Basic Degradation (0.1N NaOH) | 9966185 | 98.05 | 98.28 |
|  | 10013382 | 98.52 |  |
| H2O2 Degradation (3% Solution) | 9952769 | 97.92 | 97.95 |
|  | 9959356 | 97.98 |  |

| # 6. Limit of Detection | It is the lowest concentration of analyte in a sample that can be detected but not necessarily quantified. Limit of detection is determined by multiplying standard deviation with 3.3 and dividing by the slope of the curve. A specific calibration curve is studied using samples containing an analyte in the range of the limit of detection. The residual standard deviation of a regression line, or the standard deviation of y-intercepts of regression lines, may be used as the standard deviation. |
| --- | --- |

# Limit of Detection

| Sample | Conc. % | Conc. mg/ml | Peak Area | Standard Deviation | Slope | Conc. (ug/ml) at LOD 3.3 x std.dev /slop | Area at LOD |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 80 | 0.122 | 8032483 | 62706.37 | 69527965 | 2.98 | 200845 |
| 2 | 90 | 0.135 | 9103237 |  |  |  |  |
| 3 | 100 | 0.150 | 10122453 |  |  |  |  |
| 4 | 110 | 0.165 | 11136604 |  |  |  |  |
| 5 | 120 | 0.182 | 12235850 |  |  |  |  |

* Value of standard deviation is taken from system precision.

| # 7. Limit of Quantitation | It is the lowest concentration of analyte in a sample that can be quantified with accuracy and precision. Limit of Quantitation is determined by multiplying standard deviation with 10 and dividing by the slope of the curve. A specific calibration curve is studied using samples containing an analyte in the range of the limit of detection. The residual standard deviation of a regression line, or the standard deviation of y-intercepts of regression lines, may be used as the standard deviation. |
| --- | --- |

# Limit of Quantitation

| Sample | Conc. % | Conc. mg/ml | Peak Area | Standard Deviation | Slope | Conc. (ug/ml) at LOQ 10 x std.dev /slop | Area at LOQ |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 80 | 0.122 | 8032483 | 62706.37 | 69527965 | 9.02 | 608621 |
| 2 | 90 | 0.135 | 9103237 |  |  |  |  |
| 3 | 100 | 0.150 | 10122453 |  |  |  |  |
| 4 | 110 | 0.165 | 11136604 |  |  |  |  |
| 5 | 120 | 0.182 | 12235850 |  |  |  |  |

* Value of standard deviation is taken from system precision.

| 8. SPECIFICITY | Specificity was determined by spiking the sample with appropriate levels of excipients. Results show that the procedure is unaffected by the presence of excipients. |
| --- | --- |

# Mobile phase Interference

| # Mobile Phase | # Peak Area |
| --- | --- |
| Blank | -------- |

| Acceptance Criteria | Result |
| --- | --- |
| Method should be specific and effect of mobile phase should be negligible. | The effect of, Mobile Phase/diluent is negligible and they did not show significant effect at the elution of Avanafil It confirms that there is no interference of Mobile Phase/ diluent at appropriate retention time of Avanafil, thus the procedure is unaffected by the presence of solvents. |

| Observation/Deviation: |  |
| --- | --- |
| No deviation was observed during the validation of method of analysis for Avanafil. |  |

| Impact of Deviation: |  |
| --- | --- |
| Not Applicable |  |

| Corrective Action: |  |
| --- | --- |
| Not Applicable |  |

| Conclusion: |  |
| --- | --- |
| A comprehensive study for the Validation of the test method is performed that is being used for the determination of Avanafil Starting Material. |  |
| STATEMENT OF SUITABILITY |  |
| All the results are analyzed by the application of statistical technique i.e. Mean, Standard Deviation, Relative Standard Deviation and observed that all the results are within limits, moreover the calculated values of Coefficient Correlation, Standard Deviation & RSD are within limits. All the results of the tests carried out for the validation of the method are in complete agreement with the required limits and criteria. Method being used for the determination of Avanafil produces consistent, reproducible and reliable results therefore it is suitable for its intended purpose i.e. quantification of Avanafil. |  |